## Version Control
Performance Evaluation Reporting Template for Biomethane Conformity Assessment
Issue No.: | 1
Issue date: | 2024-10-09
Issued by: | Dr. Christopher Bamforth

## Sheet1 (hidden)
Linearity
Limit of Detection
Limit of quantification
Trueness
Repeatability
Reproducibility
Selectivity and Interference
Stability period
Uncertainty

## Reporting Tool
BiometCAP Method Validation Reporting Tool
Validation Acceptance Criteria for Determination of x - x <ENTER> <ENTER> in the presence of <ENTER> in methane by <ENTER> (Reference standard ID: <ENTER>)
Date:
Reference standard ID | <ENTER>
Analyte | <ENTER> |  | Characteristic | Critereon | Interferents | Significant interferences observed? (Y/N) | Comments
Matrix | methane |  | Selectivity | Interference between analyte and interferents | <ENTER> | - | -
Interferents | <ENTER> |  | Characteristic | Critereon | Working Range (<ENTER>) | Regression Coefficient (R2) | RSD (%) | Comments
Measurement Technique | <ENTER> |  | Linearity | Working range | x - x | - | - | -
Working Range 1 | x - x |  | Characteristic | Critereon | Units | Acheived Value (<ENTER>) | RSD (%) | Comments
Analyte Quantity Units | <ENTER> |  | LOD & LOQ | Limit of quantification | <ENTER> | - | - | -
k = 2 Expanded uncertainty target (%) | <ENTER> |  |  | Limit of detection | <ENTER> | - | - | -
|  |  | Characteristic | Critereon | Units | Target Value (%) | RSD (%) | Comments
|  |  | Trueness | Relative difference between true and measured value | % | - | - | -
|  |  | Precision | Repeatablity (μr) | % | - | - | -
|  |  |  | Intermediate precision (within lab reproducibility) (μday) | % | - | - | -
[5 empty] | Units | Target Value | Achieved value | Comments
|  |  | Uncertainty | Expanded uncertainty (U) of the method | % | <ENTER> | - | -

## Reporting Procedure
SUMMARY OF METHOD DEVELOPMENT PARAMETERS
Reporting parameter | Critereon | Definition | How to perform calculation | Section in BiometCAP protocol | Section in Eurachem guide
Selectivity | Interference | the ability of the method to determine particular analytes in mixtures or matrices without interferences from other components of similar behaviour | First, research must be performed to understand how the detector works, the types of compounds that would be detected by the method, and the types of compounds that are expected to be present in the sample. Using this knowledge, the appropriate mixture should be run using the method to compare retention times and detector sensitivity. | 4.6.2 | 6.1
Linearity | Working range | The amount fraction range in which measurements can be performed. This will depend on the set up of the analytical equipment such as sample loop and detector sensitivity. | Measure blank plus 6-10 amount fractions across the entire intended working range using either PRMS of a dynamic standard. Plot response (y-axis) against amount fraction (x-axis). Visually examine the uppper and lower boundaries, and for an initial estimate of the linear range. | 4.6.4 | 6.2
| Regression coefficient over the working range (Linearity) | When plotting the calibration curve how well does it fit a straight line. This is important when measuring a sample that falls between two points of a calibration curve. | Plot the calibration curve in excel and determine the R2 using trendline function. If the linearity is poor, reduce the working range of the method and define this as the "linear range"
LOD & LOQ | LOD (Limit of Detection) | the lowest content of analyte that can be detected by the method at a specified level of confidence (clear visible signal rather than blank/zero) | Produce low level amount fraction that is expected to be close to the LOD and perform measurmeent multiple times. Calculate LOD as 3 x s'0 (following flow chart) | 4.6.3 | 6.3
| LOQ (Limit of Quantification) | The minimum amount fraction at which a measurement result with assigned uncertainty can be made with confidence an acceptable degree of confidence. | Produce low level amount fraction that is expected to be close to the LOD and perform measuerment multiple times. It is recommended that the LOD + uLOQ is less than threshold amount fraction. Calculate LOD as 10 x s'0 (following flow chart). In this case kQ = 10 (which is commonly used), but the value can be changed by the supervisor
Trueness (bias) | Relative difference | A quantitative measure of the trueness for the method - indicating the difference between a reference value (such as traceable Primary Reference Material) and result provided by method | Produce one PRM and measure it at an amount fraction close to the threshold value that is being targeted for this method. Use Xlgenline to determine the difference between the measurement amount fraction level and the gravimetric amount fraction | 4.7.1 | 6.5
Precision | Repeatability | Measure of the variability in results when measurements are performed in a single laboratory over a short timescale. | Perform repeated measurement of the PRM at least six times in succession and calculate the relative standard deviation. | 4.7.2 | 6.6
| Intermediate precision | Measure of the variability in results when measurements are made in a single laboratory but under conditions that are more variable than repeatability conditions (e.g. on different days) | Perform measurement of the PRM at least three times using the same method and instrument but change parameters such as the analyst, day and/or timescale. Use "ANOVA:single factor" function in excel to determine the between groups variation
Uncertainty | Expanded uncertainty (k=2) | The total uncertainty the method including bias, repeatability and intermediate precision. | The total uncertainty for the measurement is calculated using the standard GC data template. | 4.8 | 6.7
Validation report |  | Final report that shall be published along with the developed method to provide details about the method, how it was validated, outcome of validation process and approved signatory | Use this validation reporting template | 4.9 | Annex A